{"text_format":"1","workbook":"xlsx","n_sheets":1,"sheets":["Shear Rate Ranges"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Shear Rate Ranges","index":0,"n_rows":21,"n_cols":17,"rows_kept":21,"cut":false,"hidden":false,"merged":[[2,10,2,11],[3,13,4,13],[5,13,11,13],[1,13,1,16]]}
["Enter viscometer parameters and sample viscosities to acquire max/min ranges"]
[null,null,null,null,null,null,null,null,null,null,null,null,null,"note: flow rates are approximate"]
["Maximum Pressure (Pa)",10000,null,"viscosity, mPa-s","min shear rate, 1/s","max shear rate, 1/s","min flow rate, ul/min","max flow rate, ul/min",null,"Chip Type","Max pressure, Pa",null,null,"Viscometer","Syringe size, uL","max flow rate (1x pump), uL/min","max flow rate (2x pump), uL/min"]
["Channel depth, um",20,null,5,{"$f":"$B$7/(D4*10/1000)*10","v":112.51125112511249},{"$f":"MIN(E4*$B$8/$B$7,$B$9*E4/G4)","v":2250.2250225022494},{"$f":"E4*60*$B$5*($B$4)^2/(6*1000*1000000)","v":0.90009000900089997},{"$f":"F4*60*$B$5*($B$4)^2/(6*1000*1000000)","v":18.001800180017995},null,"A",10000,null,null,"initium",100,1056,"---"]
["Channel width, um",2000,null,100,{"$f":"$B$7/(D5*10/1000)*10","v":5.6255625562556251},{"$f":"MIN(E5*$B$8/$B$7,$B$9*E5/G5)","v":112.51125112511249},4.500450045004499E-2,{"$f":"F5*60*$B$5*($B$4)^2/(6*1000*1000000)","v":0.90009000900089997},null,"B",40000,null,null,null,null,null,"---"]
["Min % FS",5,null,2,281.27812781278129,5625.5625562556252,2.2502250225022502,45.004500450045001,null,"C",200000,null,null,"m-VROC    hts-VROC* *(2x pump only)",100,{"$f":"1024/500*O6","v":204.8},{"$f":"2*P6","v":409.6}]
["Minimum shear stress, Pa",{"$f":"$B$3*0.01*$B$6*$B$4/8.8/2/1000*$B$5/($B$5+$B$4)","v":0.56255625562556255},null,3,187.51875187518752,3750.3750375037503,1.5001500150015001,30.003000300029999,null,"D",1000000,null,null,null,250,{"$f":"1024/500*O7","v":512},{"$f":"2*P7","v":1024}]
["Maximum shear stress, Pa",{"$f":"$B$3*$B$4/8.8/2/1000*$B$5/($B$5+$B$4)","v":11.25112511251125},null,1000,0.56255625562556255,11.25112511251125,4.5004500450045006E-3,9.000900090008998E-2,null,"E",1800000,null,null,null,500,1040,2080]
["max flow rate, uL/min",450,null,10,56.255625562556247,1125.1125112511247,0.45004500450044999,9.0009000900089973,null,null,null,null,null,null,1000,2050,4100]
[null,null,null,700,0.80365179375080364,16.073035875016071,6.4292143500064292E-3,0.12858428700012856,null,null,null,null,null,null,2500,5200,10400]
[null,null,null,700,0.80365179375080364,16.073035875016071,6.4292143500064292E-3,0.12858428700012856,null,null,null,null,null,null,10000,20500,40000]
[null,null,null,13,43.273558125043266,865.47116250086526,0.34618846500034617,6.9237693000069225,null,null,null,null,null,null,25000,50000,100000]
[null,null,null,14,40.18258968754018,803.65179375080345,0.3214607175003214,6.4292143500064274,null,null,null,null,null,"microVISC",400,450,"---"]
[null,null,null,15,37.503750375037505,750.07500750074996,0.30003000300030003,6.0006000600060005]
[null,null,null,16,35.159765976597662,703.19531953195315,0.28127812781278128,5.6255625562556251,null,null,null,null,null,null,"Chip Type","Channel depth, um","Channel Width, um"]
[null,null,null,2000,0.28127812781278128,5.6255625562556251,2.2502250225022503E-3,4.500450045004499E-2,null,null,null,null,null,null,"X02",20,2000]
[null,null,null,null,null,null,null,null,null,null,null,null,null,null,"X04",40,2000]
[null,null,null,null,null,null,null,null,null,null,null,null,null,null,"X05",50,2000]
[null,null,null,null,null,null,null,null,null,null,null,null,null,null,"X10",100,2000]
[null,null,null,null,null,null,null,null,null,null,null,null,null,null,"X20",200,2000]
[null,null,null,null,null,null,null,null,null,null,null,null,null,null,"X30",300,3000]